## Slide 1
Explainable DR Screening
Model architecture, specification and measured performance
SIH-26038 · ICDR 5-grade severity · validated zero-shot on Messidor-2
8.77 M
total parameters
grader 4.34 M + U-Net 4.43 M
2,925
training images
APTOS 2019 + IDRiD
1,744
external test images
Messidor-2, nothing fitted on it
SIGHT-THREATENING DR
97.3%
caught on Messidor-2
107 of 110 blinding eyes
100%
on the internal test split
105 of 105, none missed
Notes: Deployed model: image-only ordinal CNN, EfficientNet-B0 backbone with a CORN ordinal head. Trained on APTOS 2019 + IDRiD only. Messidor-2 is held out entirely — no training, no threshold selection, no calibration.

## Slide 2
Nine stages, every one auditable
One entry point returns a complete, JSON-serialisable result with per-stage latency
1
Geometry
FOV detect, tight crop,
square pad, resize
2
Quality gate
Ungradable images stop here
with recapture advice
3
Enhancement
Only the corrections the
gate says are needed
4
Landmarks
Optic disc 0.098 DD,
fovea 0.178 DD median, on IDRiD
5
Segmentation
Attention U-Net at 1024px,
4 lesion classes
6
Clinical features
39-dim vector: counts per
quadrant, distance to fovea
7
Grading
Ordinal CNN, temperature
+ isotonic calibrated
8
Decision
Fitted referral and urgency
thresholds, defer band
9
Explanation
Grad-CAM++ over referable
log-odds + lesion evidence
Notes: Stages 1-3 are preprocessing, 4-6 build the clinical evidence, 7-9 decide and explain. Every stage records its latency and the whole result is JSON so it can be stored, audited and replayed.

## Slide 3
The grader: ordinal, not categorical
DR severity is ordered — a 5-way softmax would treat grade 0 vs 4 as no worse than 0 vs 1
ARCHITECTURE
Backbone
EfficientNet-B0 (timm), ImageNet-pretrained
Input
512 x 512, 3 hybrid planes
Planes
CLAHE-green · Ben-Graham · L*
Embedding
1280-dim, global average pooled
Head
CORN ordinal — 4 boundary logits
Parameters
4.34 M
Selection
QWK on val, EMA weights
Sampler
sqrt-stratified, grades 3-4 to 27%
Why CORN
Each logit models a conditional: P(y > k | y > k−1). Chaining them by cumulative product gives probabilities that are monotone non-increasing by construction — the model cannot claim an eye is more likely severe than referable.
That single guarantee is what makes one score, P(grade ≥ 2), a defensible referral decision.
THE BINDING CONSTRAINT
CORN trains boundary k only on the subset with y > k−1. The grade-4-vs-3 conditional therefore sees grades 3 and 4 alone — 467 images — and it decides every severe-vs-proliferative call the system makes.
Notes: CORN: Shi, Cao & Raschka 2021. The alternative — plain 5-way softmax — has no ordering constraint, so nothing stops P(grade 4) exceeding P(grade 2). Referral would then need an ad-hoc sum over classes rather than a single calibrated score.

## Slide 4
Segmentation runs at 1024, the grader at 512
Not a tuning preference — a microaneurysm survives downsampling from 4288px as roughly four pixels
Lesion class | Dice @512 | Dice @1024 | Published range
Microaneurysm | 0.000 | 0.481 | 0.30–0.50
Haemorrhage | 0.248 | 0.539 | 0.50–0.65
Hard exudate | 0.272 | 0.572 | 0.70–0.80
Cotton-wool spot | 0.187 | 0.628 | 0.55–0.70
Attention U-Net · width 24 · 4.43 M parameters · deep supervision · focal Tversky loss (β > α, so false negatives cost more)
THE COST OF GETTING THIS WRONG
Training at 1024 and serving at 512 would hand the grader lesion counts it was never fitted on, and no metric would show it — validation reads the same cohort training produced. The resolution is recorded in the checkpoint and the deploy bundle, so the served pipeline cannot silently disagree with its own validation.
Neovascularisation is NOT assessed
No public corpus used here annotates it — verified across IDRiD (81) and DDR (532). The channel is excluded from the loss and the report says "not assessed", never "none detected".
Why it matters clinically
NV defines proliferative DR. Saying "none detected" would claim an exclusion the model never made — the difference between "we looked and it is clear" and "we never looked".
Notes: Dice 0.000 at 512px for microaneurysms is not a bug — the lesion is physically gone after downsampling. This is why the pipeline runs two resolutions and records seg_size in the deploy bundle.

## Slide 5
Every cut-point is fitted on validation data
Not one hard-coded 0.5 remains in the decision path — and none was chosen by looking at Messidor-2
Referral
on P(grade ≥ 2)
0.1999
val, sensitivity-first
—
Urgency
on P(grade ≥ 3)
0.1184
val, sensitivity-first
was: predicted grade ≥ 3
Grade boundaries
per CORN boundary
0.30 / 0.39 / 0.39 / 0.36
val, macro-recall
was: hard-coded 0.5
Lesion detection
per lesion class
0.60 – 0.95
held-out IDRiD, F1
was: blanket 0.5
Calibration
+ isotonic, tie-broken
T = 2.708
val, out-of-fold
—
Fitting the urgency cut-point took sight-threatening eyes reaching the expedited queue from 56/110 to 86/110 — with the screening sensitivity and specificity unchanged to four decimal places.
Notes: Selecting a threshold on the external set would fit a decision rule to held-out data and void the zero-shot claim. Every fit here uses the validation split only.

## Slide 6
Measured, internal and zero-shot
Messidor-2: different country, different cameras, three-specialist adjudicated grades, nothing fitted on it
Chart: 
| Sensitivity | Specificity | AUC | QWK
Internal test (n=631) | 0.986 | 0.873 | 0.988 | 0.897
Messidor-2 (n=1,744) | 0.707 | 0.922 | 0.908 | 0.649
97.3%
sight-threatening caught
zero-shot (107/110)
24.3%
of patients queued
for review
91.4%
graded within one
ICDR level
Aggregate external sensitivity is 0.707 because moderate NPDR is 76% of referable cases and Messidor-2 grades it more strictly than APTOS — a label-definition gap, not a perception failure.
Notes: Internal split is subject-disjoint from training. Both problem-statement targets — sensitivity 90%, specificity 85% — are met internally. Externally specificity holds at 0.922 while sensitivity drops, and section 3 of RESULTS.md traces that to the reference standards disagreeing on grade 2.

## Slide 7
Why the image-only arm is deployed
Four arms, each with its own temperature and threshold fitted on val, compared by paired DeLong test
Arm | AUC | Sens | Spec | QWK | Targets
cnn (deployed) | 0.9883 | 0.986 | 0.873 | 0.894 | PASS
fusion | 0.9850 | 0.989 | 0.882 | 0.882 | PASS
clinical_only | 0.9272 | 0.880 | 0.813 | 0.690 | fail
rule_based | 0.9115 | 1.000 | 0.075 | 0.000 | fail
The rule-based arm is the honest baseline
Classical ICDR criteria applied directly to counted lesions — the 4-2-1 rule, NV, exudate distance from fovea. At specificity 0.075 it refers essentially everyone, which is exactly the weakness in existing solutions the problem statement calls out. Better segmentation lifted it to 0.788, so the rules were never wrong; they were being fed poor lesion counts.
PAIRED DeLONG
vs fusion
+0.0032
p = 0.034
vs clinical_only
+0.0610
p < 1e-10
vs rule_based
+0.0768
p < 1e-11
The fusion arm was starving its own backbone: with the clinical branch zeroed it scored 0.864 where the image-only arm scored 0.956. Modality dropout restored it to 0.952.
Notes: The clinical-feature branch now powers explanation and the auditable rule trail rather than the referral score. Every referral still ships lesion counts, quadrant maps and 4-2-1 reasoning — they are simply not inputs to the number that decides referral.

## Slide 8
What it does, and what it does not
Stated because a screening tool that hides its failure modes cannot be safely deployed
VALIDATED
97.3% of blinding disease caught on unseen hardware
Both problem-statement targets met internally
Every decision cut-point fitted, none on test data
Rule trail and lesion evidence on every referral
STANDING LIMITATIONS
Printed ICDR grade is far less reliable than the referral flag — 0.333 exact on severe NPDR externally
Neovascularisation cannot be detected; PDR is never excluded on lesion evidence
Moderate NPDR at 62.3% externally — reference standards disagree
No African, East Asian or Latin American validation
Decision-support output. Not a diagnosis. Every referable and sight-threatening finding is reviewed by a qualified ophthalmologist before clinical action.
Notes: The grade-vs-flag distinction is the single most important operational message: consume the referral decision and the urgency tier, not the printed ICDR grade.